Validation of a 3D Printed Tourniquet For Hemorrhage Control in a Controlled Setting
Background
Application of a limb tourniquet is vitally important in temporizing life threatening hemorrhage following trauma in hospital and pre-hospital settings1. Currently there are no standardized testing protocols to determine the efficacy of tourniquet designs in civilian literature. In 2019, the Committee on Tactical Combat Casualty Care (CoTCCC), an arm of the U.S. Department of Defence, published a review paper on tourniquet design, including a 50-point rating scale based with both subjective and objective measures of performance including arterial occlusion scoring, speed of application, ease of use, pressure scoring and safety metrics2. This scoring system is focused on making recommendations on tourniquet designs for the U.S. Military. Throughout the years, many head-to-head tourniquet studies have been performed in controlled environments, focusing on performance metrics around arterial occlusion3–8.
Despite little innovation in design in recent decades, tourniquets can be both expensive and unobtainable during times of humanitarian crisis. Recent advancements and mass dissemination of 3D printing technologies may be able to fill this gap if these devices can create reliable and safe products. In this study we will compare the efficacy of a 3D printed tourniquet with a standard CAT Gen 7 tourniquet, focusing on performance of each device in a controlled setting.
Methods
In this trial we will as recruit EMS personnel with previous training in tourniquet application to deploy both a 3D printed tourniquet and a control CAT tourniquet to a volunteer subject’s thigh. The order of application (3D printed vs CAT) will be pre-randomized prior to application. A doppler ultrasound probe will be used to determine loss of popliteal pulse. Volunteers and subjects will be privy to doppler waveforms and feedback sounds throughout the trial to act as a surrogate for hemorrhage. EMS volunteers will then participate as the subject for the next volunteer.
Inclusion criteria will include volunteer EMS providers with prior training on tourniquet application. Exclusion criteria will include previously known vascular disease, previous lower limb surgery, previous or active DVT, ongoing infection to a limb, or open wound.
The primary outcome will be percentage of trials with successful arterial occlusion. Secondary outcomes will include time to arterial occlusion, incidence of trial failure due to mechanical failure of the tourniquet and intolerance of the procedure by the subject as the reason for trial failure. Successful arterial occlusion will be defined as 1) successful occlusion of the popliteal artery as detected by doppler ultrasound, 2) windlass mechanism locked into plastic gate and 3) arterial flow resumed at tourniquet release.
Discussion
We hypothesize that the 3D printed tourniquet is non-inferior to the CAT Gen 7 tourniquet in controlled settings. Unfortunately, placing a tourniquet in the lab is not the same as placing it in the field as situational stress can cause failure such as overtightening and damaging the tourniquet. This is discussed frequently in CoTCCC affiliated publications (ex. In ‘care under fire’ scenarios) and is likely true for EMS providers as well. We therefore hope that this controlled study will serve as the basis for more involved civilian EMS studies in the real world. Future planned studies may involve a randomized control trial of 3D printed vs standard issued tourniquets deployed by EMS colleagues in real world trauma. There are also opportunities to further develop tourniquet testing protocols to decide which designs are appropriate for pre-hospital systems .
References
1.	Walters TJ, Wenke JC, Kauvar DS, McManus JG, Holcomb JB, Baer DG. Effectiveness of self-applied tourniquets in human volunteers. Prehosp Emerg Care. 2005 Dec;9(4):416–422. PMID: 16263675
2.	Montgomery HR, Hammesfahr R, Fisher AD, Cain JS, Greydanus DJ, Butler FK, Goolsby C, Eastman AL. 2019 Recommended Limb Tourniquets in Tactical Combat Casualty Care. J Spec Oper Med Peer Rev J SOF Med Prof. 2019;19(4):27–50. PMID: 31910470
3.	Savage E, Pannell D, Payne E, O’Leary T, Tien H. Re-Evaluating the Field Tourniquet for the Canadian Forces. Mil Med. 2013 Jun 1;178(6):669–675.
4.	Kragh JF, Moore VK, Aden JK, Parsons DL, Dubick MA. Short Report Comparing Generation 6 Versus Prototype Generation 7 Combat Application Tourniquet® in a Manikin Hemorrhage Model. J Spec Oper Med Peer Rev J SOF Med Prof. 2016;16(1):14–17. PMID: 27045489
5.	Gibson R, Housler GJ, Rush SC, Aden JK, Kragh JF, Dubick MA. Preliminary Comparison of New and Established Tactical Tourniquets in a Manikin Hemorrhage Model. J Spec Oper Med Peer Rev J SOF Med Prof. 2016;16(1):29–35. PMID: 27045491
6.	Heldenberg E, Aharony S, Wolf T, Vishne T. Evaluating new types of tourniquets by the Israeli Naval special warfare unit. Disaster Mil Med. 2015;1:1. PMCID: PMC5327874
7.	Kragh JF, Newton NJ, Tan AR, Aden JK, Dubick MA. New and Established Models of Limb Tourniquet Compared in Simulated First Aid. J Spec Oper Med Peer Rev J SOF Med Prof. 2018 Summer;18(2):36–41. PMID: 29889953
8.	Glick CY, Furer MA, Glassberg CE, Sharon R, Ankory MR. Comparison of Two Tourniquets on a Mid-Thigh Model: The Israeli Silicone Stretch and Wrap Tourniquet vs. The Combat Application Tourniquet. Mil Med. 2018 Mar 1;183(suppl_1):157–161.